## Slide 1: A Novel At-Sensor Radiance Harmonization Method
Andreas Baumgartner, Claas Köhler
German Aerospace Center (DLR)

## Slide 2: Airborne Sensor Systems
HySpex VNIR-3400N + SWIR-384
N → Nyquist (sampled)
Operational since 2023
Typical hyperspectral sensor
HySpex-B
HySpex VNIR-1600 + SWIR-320me
Acquired for EnMAP preparation in 2011
Extensively calibrated
HySpex (EnMAP demonstrator)
Our advantage: we can recalibrate after each campaign

## Slide 3: Calibration Home Base (CHB)
Co-funded by ESA
Operational since 2007
Calibration of
  angular response
  spectral response
  radiometric response
  polarization
  non-linearity
  stray light
  temperature sensitivity
  …
Continuous development
Instruments: HySpex, AVIRIS-4, “EnMAP“, …

## Slide 4: Spectral Calibration Results HySpex VNIR3400-N
Individual SRF for each pixel (cubic spline model)
Complicated SRF shapes → center wavelength + resolution not well defined
Similar results for across-track response functions (ARFs)
order sorting filter edge

## Slide 5: Initial Situation
We have a very good understanding (model) of the instrument
But
  available L1 algorithms correct only smile and keystone
  higher level products do not make use of the full instrument model
How can we make use of the full instrument model
  to make data from different detector locations comparable?
  to make data from different instruments comparable?
  to reduce data complexity?
Do we really need to track the response of each individual pixel in higher level products?

## Slide 6: Imaging Equation
For simplicity only spectral case is shown
At sensor radiance
Instrument responses
Measured radiance

## Slide 7: Two Systems - Two Results
System A
System B
Ratio System A / System B

## Slide 8: Goal: Converting Data from System A to System B
Each pixel of system B is the weight and sum of all pixels of system A

## Slide 9
Baumgartner et al. (2020) doi:10.1364/oe.409626
Mathematical background in
Transformation matrix can be computed from cross-correlation matrices
Instrument properties

## Slide 10: Current HySpex VNIR-3400N Processing Settings
Parameter | Sensor A | Sensor B
Spatial columns | 3408 | 1337
Spectral channels | 700 | 295
Smile [nm] | <0.8 | 0
Keystone [mrad] | <0.14 | 0
Angular resolution [mrad] | 0.09 – 0.43 | 0.44
Spectral resolution [nm] | 2.5 – 7.5 | 4
SRF and ARF shape | Spline model (complicated) | Gaussian
smart binning
smart binning
super resolution

## Slide 11: Simulation Results
Comparison
super resolution

## Slide 12: Uncertainties
HySpex VNIR-3400N: Relative noise level change after transformation:

## Slide 13: Experimental Validation of HySpex VNIR-3400N
Measurement: Mercury lamp coupled into integrating sphere
Isolated spectral emission lines → high frequency stress test
Comparison of two spatial columns (across-track angles)

## Slide 14: Applications
Sensor A | Sensor B | Application
HySpex | ideal HySpex | harmonized L1 product smart binning
HySpex | CHIME | CHIME validation
CHIME | Sentinel-2 | cross-validation
CHIME | Sentinel-5p | wishful thinking!

## Slide 15: Conclusion
Matrix based response function harmonization is an efficient tool
  Correct several optical artifacts
  Reduce data size by smart binning
  Reduce complexity of down-stream processing
  Deliver L1 products that are easier to adopt by the community
  Works also with undersampled instruments!
Our finding have implications on the design of future sensors
  Oversampling (spectral & across-track) improves data quality
  Smile and keystone should not drive the optical design
  Knowledge of the response functions (including in-orbit changes) is of paramount importance

## Slide 16: My Answers
How does your work contribute to improving harmonization?
  I think the presentation speaks for itself
What aspect of harmonization can most benefit users?
  Uniform and comparable data products
At what level of processing is harmonization most necessary (L1B, L1C, L2A, L2B+, L3/L4)?
  Depends on instrument and application
What is more important, the harmonization on data formats or the harmonization on information content of the products?
  Data formats can be harmonized to a certain degree with wrappers
  Product harmonization is more important

## Slide 17: Geometric Calibration Results
Individual across-track response function (ARF) for each pixel
No analytical function: Cubic spline model

## Slide 18: Simulation Results Undersampled Sensor
Sensor with ~1.4 Channels / FWHM
Larger errors compared to oversampled sensor
Smaller errors compared to conventional methods
Test Scene

## Slide 19
Using Tikhonov regularization to stabilize ill-posed problem