## Slide 1
ENMAP PRODUCT VALIDATION AND
IN-ORBIT MISSION CROSS-VALIDATION
WICSIS 2024
Maximilian Brell (1, 2),
L. Guanter (3), D. Scheffler (1), K. Segl (1), S. Chabrillat (1),
M. Bachmann (4), R. Reyes (4), M. Pato (4),
E. Carmona (4), V. Krieger (5), M. Bock (5), L. La Porta (5)
and the EnMAP validation team and supporters
Helmholtz Center Potsdam, GFZ German Research Center for Geosciences, Potsdam, Germany
NEO HySpex (Norskk Elektro Optikk AS), Oslo, Norway
(3) Universitat Politècnica de València, Valencia, Spain
(4) Earth Observation Center (EOC), German Aerospace Center (DLR), Weßling, Germany
(5) Space Agency, German Aerospace Center (DLR), Bonn, Germany

## Slide 2: EnVAL: validation of EnMAP products
Data product | Field-, Image-, Model based validation |  |  | 
 | Radiance | Reflectance | Geometry | Quality
L1B | Vicarious: Inter. Sites RadCalNet Stability: PICS (Moon) SNR Scene based: Dead/bad Striping |  | Keystone MTF | SRF/Smile Anomalies
L1C | Cross-Validation (PRISMA, S2, EMIT, …) |  | Absolute Band-to-Band VNIR-to-SWIR | Anomalies
L2A |  | Ground & Water Quality mask Aerosol & WV |  | Anomalies
In-situ/
Reference data
Image based
In-situ based
Mission
Quarterly
Report

## Slide 3
L1B TOA validation
BOA REF
RadCalNet / Hypernets
In-situ spectra
AOT & WV
+ HR georeference
+ Geolayer
Propagation  (TOA RAD)
(Spectral; temporal;
view & illum.;
AOT & WV
adjustment)
Corresponding AOI/pixel; homogeneity check

## Slide 4
L1B TOA RAD validation
Complete in-orbit period
36 valid matchups
(Railroad Valley, Gobabeb, Pinnacles, Amiaz Plain, Makhtesh Ramon, Playa Zero…)

## Slide 5
L1B TOA RAD validation
3rd year OP
Only 6 valid matchups
(lack of EnMap acquisitions during field campaigns)
(Railroad Valley, Gobabeb)

## Slide 6
L1B TOA validation
Complete in-orbit period
36 valid matchups
(Railroad Valley, Gobabeb, Pinnacles, Amiaz Plain, Makhtesh Ramon, Playa Zero…)
Notes: SNR, Atmosphere, AOT retrieval, field instruments less accurate at the borders wavelength. SWIR to high. More matchups needed.

## Slide 7
L1B TOA validation
3rd year OP
6 valid matchups
(Railroad Valley, Gobabeb)

## Slide 8
L1B SNR
Complete in-orbit period
Mean SNR calculated on the basis of 93 tiles
  Criteria: VNIR > 343:1 (@495 nm & SSD 4.7 nm) SWIR > 137:1 (@2200 nm & SSD 8.4 nm)

## Slide 9
L1B spectral validation
Validation performed by
L. Guanter (UPV)
Results from 8 tiles (PV01.04.02/AV01.04.02) acquired under different atmospheric and angular conditions
  VNIR: -0.16 nm offset, ~0.35 nm peak to peak
  SWIR: 0.4 nm offset, ~0.6 nm peak-to-peak
Criteria: < 20 % linear displacement of a detector element fulfilled

## Slide 10
L1C absolute geometric &
VNIR-SWIR co-registration validation
HR Georeference
Spatial mis-registra-tion
(AROSICS)
EnMAP L1C
METRIC | VALUE [m]
RMSE_X | 5.8
RMSE_Y | 6.8
VNIR-SWIR co-registration:
Absolute geometry:
METRIC | VALUE [m]
RMSE_X | 10.2
RMSE_Y | 12.6
Great improvement!  Use up-to-date PV and AV

## Slide 11
mean spectra
L2A BOA validation (land)
ASD measurement + GPS
In-situ data provided by R. Kokaly (USGS)
Consistency; homogeneity & location

## Slide 12
L2A BOA validation (land)
Requirements: Uncertainty (RMSE)
  < 0.01 for ρ < 0.1 and
  < 0.02 for 0.1 ≤ ρ < 0.3 and
  < 0.05 for 0.3 ≤ ρ ≤ 0.6
Complete in-orbit period
47 valid matchups
(Pinnacles, Amiaz Plain, Barrax, Munich-North-Isar, Makhtesh Ramon, Camarena, EROS, TXL, Playa Zero, La Crau, Gobabeb, Railroad Valley, Nansen Ice Shelf)

## Slide 13
L2A BOA validation (land)
Requirements: Uncertainty (RMSE)
  < 0.01 for ρ < 0.1 and
  < 0.02 for 0.1 ≤ ρ < 0.3 and
  < 0.05 for 0.3 ≤ ρ ≤ 0.6
3rd year OP
3 valid matchups
(Gobabeb)

## Slide 14
L2A BOA validation (land)
Requirements: Uncertainty (RMSE)
  < 0.01 for ρ < 0.1 and
  < 0.02 for 0.1 ≤ ρ < 0.3 and
  < 0.05 for 0.3 ≤ ρ ≤ 0.6
Complete in-orbit period
47 valid matchups
(Pinnacles, Amiaz Plain, Barrax, Munich-North-Isar, Makhtesh Ramon, Camarena, EROS, TXL, Playa Zero, La Crau, Gobabeb, Railroad Valley, Nansen Ice Shelf)

## Slide 15
L2A BOA validation (land)
Requirements: Uncertainty (RMSE)
  < 0.01 for ρ < 0.1 and
  < 0.02 for 0.1 ≤ ρ < 0.3 and
  < 0.05 for 0.3 ≤ ρ ≤ 0.6
3rd year OP
3 valid matchups
(Gobabeb)

## Slide 16: EnVAL conclusions:
Data product | Field-, Image-, Model based validation |  |  | 
 | Radiance | Reflectance | Geometry | Quality
L1B | Vicarious: Inter. Sites RadCalNet Stability: PICS Moon SNR Scene based: Dead/bad Striping |  | Keystone MTF | SRF/Smile Anomalies
L1C | Cross-Validation (PRISMA, S2, EMIT, …) |  | Absolute Band-to-Band VNIR-to-SWIR | Anomalies
L2A |  | Ground & Water Quality mask Aerosol & WV |  | Anomalies
In-situ/
Reference data
Image based
In-situ based
Mission
Quarterly
Report

## Slide 17
Global statistics
L1B & L2A cross-validation
EnMAP
PRISMA
EMIT
Segmental comparison
Masking
Geometric matching
Segmenting
Filtering
Cloud detection
SZA/SAA
VZA/VAA
TOA Reflectance
Solar models
In-situ data

## Slide 18
L1B & L2A cross-validation
Filtered matchups with EnMAP:
  < 24 h timely offset
  > 1/3 off an EnMAP tile
  Cloudiness < 40%
  No off-nadir criteria
EMIT
PRISMA
EnMAP
Matchups with:

## Slide 19
L1B cross-validation
EnMAP
PRISMA
7 Matchups
41 Segments
Initial results!

## Slide 20
L1B & L2A cross-validation
Filtered matchups with EnMAP:
  < 24 h timely offset
  > 1/3 off an EnMAP tile
  Cloudiness < 40%
  No off-nadir criteria
EMIT
PRISMA
EnMAP
Matchups with:
Initiative to fill the
PRISMA and EnMAP
archives with spatiotemporal nadir matches
cross-validation
Transferfunctions e.g. for time-series densification
Notes: Work in progress:
At the moment we are working hard on the cross-validation between EnMAP EMIT, PRISMA and S2
Intersecting the missions, Masking, filtering. Everithing is segment-wise…
The goal is to generate matchups and include also in-situ data and build up a global statistik…

## Slide 21: EnVAL conclusions seed questions:
Data product | Field-, Image-, Model based validation |  |  | 
 | Radiance | Reflectance | Geometry | Quality
L1B | Vicarious: Inter. Sites RadCalNet Stability: PICS Moon SNR Scene based: Dead/bad Striping |  | Keystone MTF | SRF/Smile Anomalies
L1C | Cross-Validation (PRISMA, S2, EMIT, …) |  | Absolute Band-to-Band VNIR-to-SWIR | Anomalies
L2A |  | Ground & Water Quality mask Aerosol & WV |  | Anomalies
In-situ/
Reference data
Image based
Field based
Mission
Quarterly
Report
Independent validation is crucial for the data quality
The more matches, the better.
It is essential to represent the entire range of parameters (balancing)
Coordination of cross-matchups, including in-situ
AI for cross- CAL/VAL and harmonization

## Slide 22
In-Situ collaborators:
R. Kokaly (USGS)
B. Di Mauro (CNR)
R. Colombo (UNIMIB)
E. Ben-Dor & D. Heller (TAU)
C. Anderson (USGS)
C. Ong (CSIRO)
R. O. Green (JPL)
B. Brede (GFZ)
Niall Origo (NPL)
R. Milewski (GFZ)
A. Bialek (NPL)
C. Giardino (CNR- IREA)
M. Saberioon (GFZ)
V. Brando (CNR)
T. Schröder (CSIRO)
G. Kereszturi (Massey)
T. Painter (UCLA)
A. Kokhanovsky (GFZ)
F. Buongiorno (ASI)
…
Great community success!
Ongoing support needed
Behind the validation numbers:
Brell et al. 2024:
Assessment of EnMAP Data Quality Trough Global Product Validation Activities
Major revision RSE
Preprint available online

## Slide 23
ENMAP PRODUCT VALIDATION AND
IN-ORBIT MISSION CROSS-VALIDATION
WICSIS 2024
Maximilian Brell (1, 2),
L. Guanter (3), D. Scheffler (1), K. Segl (1), S. Chabrillat (1),
M. Bachmann (4), R. Reyes (4), M. Pato (4),
E. Carmona (4), V. Krieger (5), M. Bock (5), L. La Porta (5)
and the EnMAP validation team and supporters
Helmholtz Center Potsdam, GFZ German Research Center for Geosciences, Potsdam, Germany
NEO HySpex (Norskk Elektro Optikk AS), Oslo, Norway
(3) Universitat Politècnica de València, Valencia, Spain
(4) Earth Observation Center (EOC), German Aerospace Center (DLR), Weßling, Germany
(5) Space Agency, German Aerospace Center (DLR), Bonn, Germany